## Slide 1: Electronic System Design 3 Lecture 7.2: How to build an Instrumentation Amps
Dr Duncan Bremner

## Slide 2: Difference Amplifier
First shot at building an INA:
Assume perfect opamp Vn ==Vp
Vp
Vn

## Slide 3: Difference Amplifier (2)
Works as a gain of 1 instrumentation amplifier.
Can we have gain >1?

## Slide 4: Difference Amplifier (3)
Circuit rejects common-mode signals because it is symmetrical.
Problem arises due to resistor tolerance which makes the circuit unsymmetrical.
Worst case error when sides are maximally unbalanced:
(Assume all Resistors are +/- 1%)
Gain 2%
too small
Gain 2%
too big

## Slide 5: Difference Amplifier (4)
Easier to calculate if we amalgamate the errors in each side:
is the error in R2 (in Ohms!). In the above case the TOTAL error can be up to 2% for 1% resistors
(i.e. 1% error in R1 plus 1% error in R2)

## Slide 6: Difference Amplifier (5)
Assume differential gain ≈ 1, so R2 = R1

## Slide 7: Difference Amplifier (6)
Calculate Common-Mode and Differential gain separately:

## Slide 8: Difference Amplifier (7)
Differential gain ≈ 1

## Slide 9: Difference Amplifier (8)
Ohms
Ohms

## Slide 10: Difference Amplifier (9)
Resistor mismatch gives a small error in differential gain (trivial)
e.g. 1% resistors, CMRR = 50 (often expressed in dB “34dB”)
0.1% resistors, CMRR= 500 (54dB)

## Slide 11: Difference Amplifier (10)
But Sensitive to impedance mismatch at input
Can use precision resistors…...
Stray errors due to:-
  • Contact resistance, PCB tracks
  • Thermocouple wire resistances (e.g. Cu / Constantan)
  • Oxidation, bad soldering
Gets worse with age

## Slide 12: Difference Amplifier (10)
Can fix source impedance variation by buffering
But still limited by resistor matching. Best case ~ 0.01% (discrete)
Monolithic ~0.001% (Texas INA105, Analog Devices AMP03)

## Slide 13: Difference amplifier (11)
Typical example: bridge amplifier (scales / load cell)
Need common-mode rejection ratio of 54dB (0.1%) to see anything
Need common-mode rejection ratio of 114dB to resolve 1g in 1kg
Expensive resistors then trim. Calibrate often.
Ugh!
(Full scale)

## Slide 14: Typical Load Cell Spec

## Slide 15
Thank you
谢谢